## Slide 1: Comparison of Online and Offline ED-XRF Techniques for Atmospheric PM10 Measurement
G. Deluca1, E. Bloise1, D. Cesari1, A. Dinoi1, A. Pennetta1, S. Potì1-2, P. Semeraro1, F. Unga1, D. Contini1
1 Institute of Atmospheric Sciences and Climate, ISAC-CNR, Lecce, 73100, Italy
2 Department of Engineering for Innovation, University of Salento, Lecce, 73100, Italy

## Slide 2: INTRODUCTION
EXPERIMENT | ED-XRF (offline) | XACT (online)
SITE | ECO (*) (LAB-BASED XRF ANALYSIS) | ECO (IN SITU)
COLLECTED FILTERS | (*)samples was collected on filters using a dual-channel sampler at 2.3 m3 h−1 | instrument automatically samples and analyses spots
SUBSTRATE | quartz | teflon
PM | PM 10 | 
TIME RESOLUTION | 24 h | 3h
# OF ELEMENT | 19 elements from Al to Pb | 36 elements from Al to Bi
(*) Unga, F., et al., Determination of aerosol composition by ED-XRF on Teflon and quartz substrates: potentialities and limits. Aerosol Research, 2025. 3(2): p. 405-415

## Slide 3: METHODS
METHOD | ED-XRF (offline) | XACT (online)
TIME | From December 2024 to June 2025 (76 sample) | 
ELABORATION DATA | Data < MDL replaced with 0.5 x MDL | 3-h samples were averaged to 24-h to match daily filters. less than 6 valid 3-hourly values available. Data < MDL replaced with 0.5 x MDL
INTERCOMPARISON | 13 elements with concentrations consistently above the MDL for both online and offline techniques: (S, Cl, K, Ca, Ti, Cr, Mn, Fe, Cu, Zn, Br, Sr, Pb) | 
STATISTICAL ANALYSIS | Linear regression, R2 | 

## Slide 4
ELEMENT | LINEAR REGRESSION (SLOPE) | R2
SULPHUR | 0.75 | 0.79
CHLORINE | 0.62 | 0.69
POTASSIUM | 0.75 | 0.85
CALCIUM | 0.89 | 0.97
TITANIUM | 1.13 | 0.99
CHROMIUM | 0.69 | 0.91
MANGANESE | 1.01 | 0.92
IRON | 0.77 | 0.99
COPPER | 0.96 | 0.82
ZINC | 1.05 | 0.84
BROMIUM | 0.68 | 0.60
STRONTIUM | 1.02 | 0.95
LEAD | 1.05 | 0.96
RESULTS: intercomparison of Xact and ED-XRF

## Slide 5: RESULTS: intercomparison of Xact and ED-XRF
Chart: Pb
Chart: Sr
Chart: K
Chart: Ca
Chart: S
Chart: Mn

## Slide 6: TIME RESOLUTION OF XACT BOUNDARY LAYER FIREWORKS 1 JANUARY
Chart: fireworks
| 45657.020833333336 | 45657.125 | 45657.25 | 45657.375 | 45657.5 | 45657.625 | 45657.75 | 45657.875 | 45658.020833333336 | 45658.125 | 45658.25 | 45658.375 | 45658.5 | 45658.625 | 45658.75 | 45658.875 | 45659.020833333336
ALUMINIUM | 532.4001 | 492.3895 | 555.5508 | 395.7785 | 278.7986 | 374.4613 | 476.7131 | 592.1375 | 3252.3324 | 2128.1503 | 610.9676 | 332.3256 | 327.1362 | 336.3074 | 664.5178 | 501.4254 | 646.6573
SULPHUR | 659.9471 | 628.8281 | 572.6427 | 520.719 | 375.3595 | 411.5693 | 644.5623 | 825.5012 | 4166.7888 | 2795.9434 | 758.6694 | 411.2028 | 304.0567 | 340.8319 | 445.9538 | 751.4527 | 830.1846
BISMUTH | 0 | 0 | 0 | 0 | 0 | 0 | 0.5619 | 1.6983 | 231.3923 | 126.1398 | 19.8294 | 3.3969 | 0.5355 | 1.2488 | 1.1566 | 0.5984 | 0.7634
BARIUM | 4.223 | 2.1859 | 1.0543 | 0.7444 | 0.127 | 2.8432 | 6.9496 | 8.9446 | 837.1526 | 460.4524 | 72.4804 | 15.4056 | 3.4487 | 7.6587 | 21.4366 | 6.5181 | 6.6309
STRONTIUM | 0.3308 | 0.2708 | 0.0658 | 0.1885 | 0.3378 | 0.3034 | 0.7213 | 2.7542 | 259.22 | 174.6176 | 24.8441 | 5.0348 | 1.5927 | 2.4694 | 2.6896 | 1.0188 | 1.6367

## Slide 7
CONCLUSION
Results obtained with the two XRF techniques were compared in order to verify the goodness and reliability of the analyses of XAct;
Different substrates were used for comparison (Teflon and quartz) and 13 elements were compared;
The concentrations of the elements compared show overall good correlation (best correlations for Ti, Mn, Sr, Pb);
Participation in the MITRAP project.
FUTURE OBJECTIVES
Investigating possible underestimation of ED-XRF during high pollution events;
Intercomparison in which the same filters will be analysed by both system.