## Slide 1: VHEE generation and transport: Experimental Work
M. Salvadori, F. Avella, F.Baffigi, G. Bandini, G. Benincasa, F. Brandi, M. Ezzat, A. Fregosi, L. Fulgentini, D. Gregocki, P. Koester, L. Labate, D. Palla, M. C. V. Panaino, S. Piccinini, S. G. Vlachos & L. A. Gizzi.
Milestone 1.1 and 1.6

## Slide 2
Radiation therapy is used to treat around 60% of all cancer cases worldwide
Deliver lethal dose of radiation to tumor cells
Sparing healthy tissues

## Slide 3
Radiation therapy is used to treat around 60% of all cancer cases worldwide
Deliver lethal dose of radiation to tumor cells
Sparing healthy tissues
L. Whitmore et al., Scientific Reports 11,14013 (2021)
Photons
ease of production, routinely applied
Not suitable for deep seated tumors
High entrance dose
Protons
Dose deposition Bragg Peak,
Ok for deep-seated tumors
Cost and facility availability,
highly affected by inhomogeneities
Carbon ions
Even sharper Bragg peak
high relative biological effectiveness
Cost and facility availability, affected by inhom.
VHEE 50 MeV E 350 MeV
High entrance and exit dose
Ok for deep-seated tumors
Sharp beam penumbra
Low sensitivity to inhomogeneities
Electron 5 MeV E 20 MeV
Not suitable for deep seated tumors
High entrance dose
High lateral spread
Sharp fall off after 5cm (IORT, TSEBT)
Dose deposition profiles in tissue equivalent media

## Slide 4
L. Whitmore et al., Scientific Reports 11,14013 (2021)
Photons
ease of production, routinely applied
Not suitable for deep seated tumors
High entrance dose
Protons
Dose deposition Bragg Peak,
Ok for deep-seated tumors
Cost and facility availability,
highly affected by inhomogeneities
Carbon ions
Even sharper Bragg peak
high relative biological effectiveness
Cost and facility availability, affected by inhom.
VHEE 50 MeV E 350 MeV
High entrance and exit dose
Ok for deep-seated tumors
Sharp beam penumbra
Low sensitivity to inhomogeneities
Electron 5 MeV E 20 MeV
Not suitable for deep seated tumors
High entrance dose
High lateral spread
Sharp fall off after 5cm (IORT, TSEBT)
Radiation therapy is used to treat around 60% of all cancer cases worldwide
Deliver lethal dose of radiation to tumor cells
Sparing healthy tissues
Dose deposition profiles in tissue equivalent media

## Slide 5
FLASH - RT
Delivering high doses in an extremely short irradiation time (Dose rate >40 Gy/s, 1 Gy per pulse…)
Damages to cancer cells equivalent to conventional RT but reduced toxicity for healthy tissues!
VHEE
FLASH
VHEE
FLASH
VHEE FLASH-RT
dosimetric characteristics of VHEE
+
Sparing capability of FLASH

## Slide 6
An emptying zone is created:
The plasma bubble.
T. Tajima and J. M. Dawson Physical Reviev Letters 43, 267 (1979)
High intensity laser pulse exerts a ponderomotive force on plasma charges.
Electrons, being lighter than ions, experience a larger displacement from their rest position.
The displaced electrons oscillate around their rest position at the plasma frequency.
As the driver pulse travels in the plasma the negative charges accumulate at the back of the bubble.
Charged particles moving in the same direction as the laser pulse are accelerated!
The electric field generated by the charge separation exhibits a longitudinal component
F. Avella
High intensity laser,
Gaseous target
Plasma Generation

## Slide 7
Relatore:
Evento:
Nome relatore
Nome evento
USER CONTROL
ROOM
TESTING AND PROTOTYPING
LASER FRONT END
10 TW, 10Hz
POWER AMPLIFIER Up to 240 TW
SHIELDED TARGET AREA FOR PARTICLE ACCELERATION
NEW BEAMLINE FOR PRE-CLINICAL STUDIES
LASER DEV. LAB
NEW HIGH DOSE UNDERGROUND BUNKER
LASER CAPABILITIES:
• 240 TW, Ti:Sa, up to 5 Hz, 27 fs;
• 1kHz, >20 mJ, Ti:Sa + OPA (commissioning in progress)
• 100 Hz, >1J, TiSA (procurement in progress)

## Slide 8
Circular aperture 2 mm
ILIL laser 240 TW system
1018 Wcm-2
3 J in 27 fs
800 nm
e-
Supersonic flow of Helium Gas with 1% Nitrogen
LWFA Ionization Injection
M. Chen et al., Physics of Plasmas 19, 033101 (2012)
Output window
Beam Size and Pointing
Beam Charge
Beam Spectra
Dose and DPP
Operative electron density
Interaction Chamber

## Slide 9
S. Vlachos paper in preparation
D. Gregocki paper in preparation
Beam Charge
Beam Spectra
Dose and DPP
S. Piccinini
e-
Integrating Current Transformer (ICT)
ICT controller
Scope
Non destructive
On line charge
estimate
Typical beam charges:
Lanex
Mirror
Camera
Magnetic
dipole
e-
collimator
Typical peak energies: MeV
recorded spectrum image
RCF stack
e-
Longitudianl profile
Transverse profile
Cumulating 120 consecutive shots

## Slide 10
S. Vlachos paper in preparation
D. Gregocki paper in preparation
Beam Charge
Beam Spectra
Dose and DPP
S. Piccinini
e-
Integrating Current Transformer (ICT)
ICT controller
Scope
Non destructive
On line charge
estimate
Typical beam charges:
Lanex
Mirror
Camera
Magnetic
dipole
e-
collimator
Typical peak energies: MeV
recorded spectrum image
RCF stack
e-
Longitudianl profile
Transverse profile
Cumulating 120 consecutive shots
Follow S. Vlachos Talk !
Follow D. Gregocki Talk !
Follow S. Piccinini Talk !

## Slide 11
Beam Size and Pointing
D. Gregocki
Distance Lanex screen to source: 57 cm.
Typical signal acquired during experiments
Typical measured beam size: 1.15 0.90 cm
Typical fluctuation around centroid:

## Slide 12
C. M. V. Panaino L. Labate
Magnetic Beamline main effects:
Enhance pointing stability
Energy selector
Reducing beam size
Beam Size and Pointing with MBL during experiments
Typical beam size at 57 cm from source: 0.60 0.44 cm
Fluctuations around centroid 0.46 mm
e-
Geant4
Geant4

## Slide 13
C. M. V. Panaino L. Labate
Magnetic Beamline main effects:
Enhance pointing stability
Energy selector
Reducing beam size
Beam Size and Pointing with MBL during experiments
Typical beam size at 57 cm from source: 0.60 0.44 cm
Fluctuations around centroid 0.46 mm
e-
Geant4
Geant4
Further investigated in LPA-STAR cascading calls

## Slide 14
Relatore:
Nome relatore
Thank you for your atttention!
Very High Energy Electron have an extraordinary potential for the application in radiation therapy. The possibility of deliver them in short-temporal bunches characterized by an ultra-high dose rate pave the way to a VHEE FLASH-RT combining the benefits of both.
Laser-driven electron acceleration can be used to drive such electron bunches.
A brief overview of the laser-driven electron beamline at ILIL has been given providing its main features. Electron spectra peaked beyond 150 MeV, charge per bunch of ≈ 150 pC delivered in tens of fs are routinely obtained.
The VHEE beamline has already been used to irradiate peripherical human blood samples. The radiation damage has been evaluated through the micronuclei and telomere shortening biological endpoints. A new set of measurements on cancer cell is scheduled with an improved irradiation system which is currently being developed.
A new gas system is currently being tested as a preparatory phase for enabling high-repetition rate workflow.
A modular and cost-effective magnetic beamline has been designed. Its capability of reducing beam size and enhancing pointing stability has been experimentally demonstrated.
Follow A. Borghini Talk !

## Slide 15
Relatore:
Evento:
Nome relatore
Nome evento
This work was supported by the PNRR MUR project ECS00000017-THE “Tuscany Health Ecosystem”, funded by the European Union –Next Generation EU.
Advanced Photon Sources
PNRR MUR project IR0000016-I-PHOQS, funded by the European Union –Next Generation EU
EU Horizon IFAST, under Grant Agreement No.101004730
EuPRAXIA Advanced Photon Sources - EuAps (IR0000030, CUP I93C21000160006).